Clean steam generator纯蒸汽发生器

Intended use: Designed to generate clean steam for sterilization of pipelines and capacitive equipment as well as to provide the work of sterilizers in pharmaceutical production.

用途：产生清洁蒸汽用于消毒管线和电容设备，同时提供医药生产的消毒工作。

Capacity:产能

| Feed steam pressure, bar 蒸汽压力 | Capacity, kg/h小时/公斤 |
| --- | --- |
| 4 - 6 | 1000 |

Equipment:设备

- FW injection pump manufactured by Grundfos，

由Grundfos生产的FW注射泵

- system of measurement and adjustment of the level in columns

测量和调整塔器水平的系统。

- industrial steam feed control system

工业蒸汽供给控制系统

- clean steam feed control system

干净蒸汽供给系统。

- clean steam flowmetering system

干净蒸汽流量系统

- particulate matters filtering, separation and condensate separation system at the inlet of industrial steam to steam generator

在工业蒸汽进入蒸汽发生器入口的细小物体的过滤器，分离和冷凝分离系统

- valves/control produced by Gemu or Burkert

由Gemu or Burkert生产的阀

- steam and condensate components produced by Spirax sarco.

- 由Spirax sarco生产的蒸汽和冷凝器件。

Materials: stainless steel, all feed water and clean steam-contact parts should be manufactured of materials 1.4435/1.4404 (according to USA AISI 316 L or AISI 316 Ti standards), other parts – of material 1.4301 (according to USA AISI 304 standard).

材料： 不锈钢：所有注入的水和清洁蒸汽接触的部件按USA AISI 316 L or AISI 316 Ti标准，其他部件按USA AISI 304标准。

Welded joints shall be carried out by means of inert gas welding. All units of the installation should be accessible for inspection and easily dismantled.

用惰性气体焊接。 所有安装的单位可以拆解检验，并容易拆卸。

Surfaces: 表面抛光

- polished inside	内部Rа < 0.6 µm

- abraded outside外部	Rа < 0.9 µm

-welded joints of the machine units in working areas shall be polished to the base metal state. Shell metal – polished stainless steel.

设备焊接部件在工作区域应当抛光到基础状态。 外壳抛光不锈钢。

Basic control and automation functions:基础控制和自动功能

-- PLC controller and sensor control panel on the basis of Simatic Siemens

用西门子Simatic PLC和传感控制操控台
-- level monitoring of source water in the first distiller column

第一个蒸馏塔水位监视
-- steam pressure and heating control

蒸汽压力和加热控制
-- compressed air pressure control

压缩空气压力控制。
-- clean steam pressure control with an output of (4-20) mA for external tabbing

4-20mA输出的干净蒸汽压力控制用于外工作台。
-- permanent control of clean steam parameters with the help of clean steam condensate conductivity, with possibility of registration of inlet (FW flow) and outlet (electric conductivity of condensate and clean steam pressure) parameters from (4-20) mA output for external tabbing of digital monitor

在清洁蒸汽冷凝传导的帮助下，（冷凝和清洁蒸汽压力的电传导）出口参数
-- respiratory filter and valves for air bleeding from steam generator

通过蒸汽发生器流过的空气呼吸过滤器和阀门

- the printout or digital version of PLC controlling program with comments

PLC controlling program with comments can be printout

可以打印PLC控制程序内容

-- the possibility of SCADA Ethernet interface connection to the system 
SCADA 以太网界面连接系统

Documentation:

- pressure vessel certificate translated into English and executed on form provided in Appendix 3 (Guidelines on Vessels), with drawings, connection diagram and stress analysis of the vessel and its elements (the vessel should be made with compliance to the regulations of State monitoring of work safety) and weld crack detection;

压力容器证书英文版，带图和连接图表和容器压力分析和原理（所做容器应该和国家安全监察规定相符。

- certificate for connecting pipelines;连接管道证书。

- documents on performed inspection and testing while production of the pressure vessel;

压力容器生产时的检查和测试文件

- instruction on equipment assembly and disassembly (installation instruction);安装说明

- equipment mounting instruction;设备装配说明。

- operating instruction;操作说明

- equipment cleaning instruction;清洁说明

- maintenance instruction (Maintenance Schedule, servicing and full repair works);

维修保养说明（维修计划，服务和全部的修理工作。

- problem checklist and troubleshooting methods;故障处理表

- certificates for component parts;部件的合格证书。

- certification for the valve and instruction on the safety valve adjustment with indication of the valve flow coefficient for compressible and non-compressible media as well as installation area;

- specifications for all equipment elements;

- certificates for materials and control valves;

- up-to-date sheets on gauging, validation and calibration of measuring equipment;

- software certificate of conformity;

- documentation in English .

Description of produced clean steam:

| Value name | Permissible rates |
| --- | --- |
| Temperature | (123 ( 140)(С |
| Specific conductivity | ≤1,1 µS/cm at 20°С |
| Description | Water steam condensate – transparent colourless liquid, odorless and tasteless. |
| Chlorides | Not more than 0.00005 % (0,5 ррm). |
| Nitrates | Not more than 0.00002 % (0,2 ррm). |
| рН | 5.0 – 7.0. |
| Heavy metals | Not more than 0.00001 % (0.1 ррm). |
| Oxidants | Pale-pink colour should not disappear |
| Bacterial endotoxins | Not more than 0.25 МЕ in 1 ml |
| Total organic carbon level (TOC level) | Up to 500 ppb |

Feed water characteristics – the possibility of the steam generator running with the feed water with specific conductivity of up to 4.3 μS/cm – 1 at 20°С.

Characteristics of industrial steam at the steam generator inlet – pressure of 4 – 6 bar,

Operating costs: Please, provide the spare and wearing parts list as well as the materials subject to planned change during equipment operation, their price, change intervals and manning level necessary for maintenance schedule works performance. The information should be time-ranged: from the day of commissioning until the full repairs term.

消毒柜

Destination: sterilizer for sterilization injectable drugs in glass ampoules with the possibility to control hermetic of ampoules by a crash-test or vacuum test.

Productivity: working volume of sterilization chamber should provide simultaneous sterilization of 98 cassettes with ampoules size 300*300*100 mm, with a total cycle time (sterilization + vacuum check + drying) - not more than 3 hours (or smaller size chambers with a longer cycle time)
Volume of chamber   -               1400 -  1500 liters.
Volume of casing (jacket)  -     200   -  250 liters.

Number of ampoules 10 ml in a single download - 32 000 pcs

Number of ampoules  5 ml in a single download – 47 000 pcs.

Number of ampoules  2 ml in a single download – 80 000 pcs.

Number of ampoules  1 ml in a single download – 80 000 pcs.

Working conditions:

Working pressure:          -1 + 2.5 bar.

Casing:                          -1  + 5,0 bar.

Working temperature from 70°С to 137°С at the maximum allowable pressure of 2.5 bar. According to DSTU ISO 11134:2003 “Sterilization of medical products.

Requirements for conformity assessment and streaming control. Industrial wet heat sterilization.” Sterilization effect F0 must be not less than  15 minutes . Time of sterilization must be not longer than 3 hours (time includes sterilization + checking for leaks + drying ).

Options:
- sterilization with pure steam;
- sterilization with superheated water;
- sterilization with air-steam mixture;

The supplier must specify the benefits of sterilization type, which he suggested, with the ratio of : price/quality, cost of equipment / operating cost, time of the process, using efficiency of chambers volume compared with other types of sterilization.

Provide:

- horizontal, two-door (passing through) sterilizer type

- temperature and pressure management and control during the process( in chamber and casing)

- security control equipment in chamber and casing (jacket)

- presence of a flexible sensor for monitoring temperature in the product

- pall filter for  air supply into the chamber

- the possibility of an emergency stop

- regulation of sterilization is carried out by  pressure and temperature with conversion t F0

- equipped with doors on sliding bearings

- possibility to adjust sterilizer height

- heat-isolating outer casing is made of steel AISI 304 with high-quality isolation

- cover must be easily removable

- the ability to install doors for access to the technical area

- device to measure the amount of air and non-condensed gas in the sterilization cycle for the registration of possible chamber leaks, not identified in the performance of the vacuum test.

- the required number of temperature sensors

- Control panel with «touch-screen» display with the possibility of an ergonomically optimal installation.

- the possibility of changes in process parameters (temperature, time, pressure, depth of vacuum);

- preservation of the various programs in memory

- registration of process parameters: temperature, time, pressure, storage, program execution memory.

- thermal printer for printing checks with the program parameters performance. Check shows: the identification number of sterilizer, date and time of sterilization, temperature and pressure changes over time, the process cycle.

- Ability to print saved processes of previously performed sterilization.

- the use of vacuum test and the Dick Bove-test programs.

- manual release of pressure  in case of emergency

- door locking system when the chamber is under pressure and if there is presence of  foreign object near the doors.

Provide equipment:

Vacuum liquid-packed ring pump.

Blow-off  valve in chamber and jacket.

All pipeline connections and isolation valves must be «Tri-clamp» type.

Type of attachment «Tri-clamp» for sterilization connection to   workshop utilities.

SIEMENS Software.

Trolleys with directing or perforated shelves for transporting, loading, uploading, and sterilization of ampoules in cassettes, in an amount not less than 2 pieces.

Shut-off-and-regulating fixtures of average and membranous GEMU type.

Material: chromium-nickel steel. All details which connects with product must be made of materials according to standard USA AISI 316 L, the rest of materials  according to standard USA AISI 304. Isolation must provide the temperature of external jacket surface not more then 45°С. Seal material  PTFE or EPDM

Surfaces:

- polished inside  Rа< 0,8 µm

- field welds of machine units in working areas must be etched and polished to the condition of main metal. Material of jacket – polished stainless steel.

Steam autoclave for clothes and options

Intended use:

The autoclave shall provide sterilization and drying of:

- the packs with rubber stoppers for infusion solution vials;

- the packs with aluminum caps for infusion solution vials;

- clean-room clothing assemblies;

- adjustable dosing units of infusion and injection solutions filling machines;

- silicone hoses;

- Pall sterile filters;

- cassettes for ampoules dimensioned 300х300х100 mm.

The autoclave shall provide drying of rubber stoppers and clothing. Residual humidity of rubber stoppers < 0.2%,

Effective volume:

- Effective volume of chamber- approximately 1.5 m3

- the autoclave should meet the following standards requirements: ISO 11134:1994, EN 285, HTM 2010 and National Standards of Ukraine ISO “Sterilization of medical devices. The requirements for verification of conformity and inspection control. Industrial sterilization with moist heat.”

- maximum autoclave width – 2000 mm. (1800 mm preferable)

- maximum autoclave length – 2500 mm.

The supplier should indicate the advantages of the autoclave type which is being offered in the sense of: price/quality relationship, equipment cost/operating cost, sterilization process time, efficiency of the chamber volume usage comparing to other sterilization types.

Design version – clean steam sterilization.

The following should be provided for:

- two-door (through-type) sterilizer, horizontal

- control system and temperature and pressure control during the process (in the chamber and jacket)

- safety system measuring instruments in the chamber and jacket

- flexible transmitter for temperature control in the product

- Pall filter for air supply to the chamber

- emergency stop of the process

- the sterilization is adjusted on the basis of pressure and temperature corrected to F0

- fitting with doors with sliding guides

- sterilizer height adjustment option

- surface heat-insulating cover made of steel AISI 304 with high-quality insulation.

- easy-removable coat

- possibility of installation of doors to access the processing area

- device for estimation of the air and uncondensed gases amount during sterilization cycle to register possible chamber leakages not discovered while vacuum testing

- required quantity of temperature transmitters

- ergonomically-designed control panel with touch screen display

- process parameters measuring option (temperature, time, pressure and vacuum depth)

- memory retention of various programs

- registration of process running according to these parameters: temperature, time, pressure, memory retention of completed programs.

- thermal printer for printing of check card with completed program parameters. On the check card such information should be printed: sterilizer identification code, date and time of the processing, temperature, pressure time changes, cycle of  process.

- possibility of previously run sterilization processes printout.

- application of vacuum test and Bowie-Dick test programs

- manual pressure relief in case of emergency

- door interlocking when the chamber is under pressure or in case of foreign objects in the door area.

Materials: Chrome-nickel steel, all product-contact parts should be manufactured of materials 1.4435/1.4404 (according to USA AISI 316 L or AISI 316 Ti standards), other parts – of material 1.4301 (according to USA AISI 304 standard). The insulation should provide external surface temperature of insulating shell of not more than 45°С.

Surfaces:

- polished inside		Rа < 0.9

- abraded outside		Rа < 1.2

- welded joints of sterilization chamber shall be polished to the base metal state. Shell metal – polished stainless steel.

The following equipping shall be provided for:

- Safety valves in the chamber and jacket. Autoclave is controlled with the control unit with programmable memory for registering of minimum 5 sterilization cycles; the menu of control unit display shall be Russified;

- Inner trolley with shelves for sterilized material and two floor vehicles for feed and discharge;

- Water-ring vacuum pump to support absolute pressure of ≤ 4 KPa, produced by Lema MEX125;

- Pressure safety system, mechanic and electronic door interlocking system;

- The 2-door design shall provide for the opening of only one door, whereas another door shall be hermetically sealed;

- Sterilization effect F0 shall continue for not less than 15 min, in compliance with European Pharmacopoeia requirements;

- Condensate shall be discharged to the export header directly.

All pipe, valves and fittings connections shall be of Tri-clamp type.

The connection of sterilizer to workshop energy carriers shall also be Tri-clamp.

Software type – SIEMENS.

Shutoff and control valves shall be of seat and diaphragm GEMU type.

Steam and condensate components produced by Spirax sarco.

Door seal drive shall be air-driven.

Appendix 1*.

Operating conditions:

- maximum working pressure inside the chamber: 1 – 3 bar (differential)

- jacket maximum working pressure:	0 – 3 bar (differential)

- chamber rated pressure: 			-1 – 3,5 bar (differential)

- jacket rated pressure:				0 – 3,5 bar (differential)

- rated temperature:				1400С

Documentation:

- pressure vessel certificate translated into Ukrainian or Russian and executed on form provided in Appendix 3 (Guidelines on Vessels), with drawings, connection diagram and stress analysis of the vessel and its elements (the vessel should be made with compliance to the regulations of State monitoring of work safety) and weld crack detection;

- certificate for connecting pipelines;

- documents on performed inspection and testing while production of the pressure vessel;

- the finding (certificate of conformity) of Kiev Technical Centre on conformity of the autoclave to the working safety requirements of Ukraine;

- installation and mounting instruction;

- operating instruction;

- equipment cleaning instruction;

- maintenance instruction (Maintenance Schedule, servicing and full repair works);

- problem checklist and troubleshooting methods;

- certificates for component parts;

- certificates for the valves and instruction on the safety valve adjustment with indication of the valve flow coefficient for compressible and non-compressible media as well as installation area;

- certificate for springs used in the safety valves;

- specifications for all equipment elements;

- certificates for materials the autoclave, valves and controls are made of.

- up-to-date sheets on gauging, validation and calibration of measuring equipment, see Appendix 1*;

- two temperature transmitters (registering one and controlling one);

- two pressure transmitters (registering one and controlling one);

- two pressure gauges in the chamber (one per each outlet), jacket-covered;

- documentation in the Russian or Ukrainian language.

Operating costs: Please, provide the spare and wearing parts list as well as the materials subject to planned change during equipment operation, their price, change intervals and manning level necessary for maintenance schedule works performance. The information should be time-ranged: from the day of commissioning until the full repairs term.

Existing utilities:

- Compressed air 			 					6 bar

- Steam the condensate of which corresponds with WFI 		3 ± 0.2 bar.

- Power supply			 					380 V, 50 Hz.

- Water					 				2-3 bar.

Input resources: Please, identify the quality and quantity of energy materials which should be supplied to your equipment to provide its appropriate and qualitative functioning (steam, condensate bleeding, water supply to vacuum pump, power supply).

Appendix 1*

- The name plate in Russian or Ukrainian should be placed on the autoclave, stating:

- the name and address of the manufacturer, serial number or other identification of the equipment,

- rated pressure inside the chamber and maximum working temperature,

- pressure inside the chamber jacket,

- surveillance inspector’s mark and vessel identification label,

- production date of the vessel.

2.	The recording device and control system shall be independent and designed in such a way to sound the alarm in case the differences between the controlled and recorded variables exceed the allowable limits.

Measuring instruments:

- The temperature control device should be:

- digital or analogue,

- with accuracy of ±1% in the range of scale of 50 - 150 0С,

- adjusted with accuracy of (±0,5) 0С of sterilization temperature value,

- with a transmitter protected against destruction.

- The pressure control device should be:

- digital or analogue,

- with accuracy of ±1.6% in the range of scale of 0 - 5 bar,

- with a transmitter protected against destruction.

3)	the timer should be with accuracy of not less than ±1% for time periods of more than 5 minutes and not less than ±2,5% - for time periods of less than 5 minutes.

3.	Calibration of instruments (2 temperature and 2 pressure transmitters).

The calibration program should be documented. The calibration program shall make reference to regulatory requirements, written calibration procedures and directives of instrument precision.

Instrument calibration should be carried out within the range of 2 0С of sterilization value (1210С and 1320С).

- The possibility of independent selection of values of sterilization and drying (the temperature and time of sterilization), and the name of product being sterilized.

Therewith pressure inside the chamber (to obtain preset temperature) should be set automatically (on the basis of the factory tests and calculations).

5.	There should be automatic changing of the steam pressure inside the jacket.

Note: For uniform distribution of temperatures in the autoclave chamber (not more than 20С) it is necessary that the defined pressure (temperature) inside the chamber corresponds to the defined pressure (temperature) inside the jacket.

Consequently after the desired temperature has been set, the pressure inside the jacket (or chamber) changes automatically (on the basis of the factory tests and calculations).

- The equipping with safety system for protection from entering to autoclave menu. The access to menu should be protected with password.

There should be a different passwords for each operator for further indication of the code (or surname) of the operator running the process on the printout.

- Autoclave printer should be printing during the whole sterilization or drying process (up to 300 minutes) with the option of spacing change.

- The printouts shall contain the following information:

- the name of product being sterilized,

- the preset parameters of preliminary air removal, sterilization and drying (compulsory: time, temperature and pressure),

- autoclave name and its serial number,

- the name or code of the operator running the process (this could be realized through different passwords)

- the date and time of the start of sterilization,

- the date and time of the end of sterilization,

- the date and time of the start of drying,

- the date and time of the end of drying,

- obtained value of sterilization effect F0.

- current temperature and pressure values inside the chamber at a given time.

- Software: the manufacture should supply the approved protocol of software validation which includes the reference to the parameters and standards of the carried out software validation.

- While vacuum relief it is necessary that the air ingoing into the chamber pass through the filter with the efficiency of not less than 99.5 % of particles with the size of more than 0.3 μm.

There should be a calibration certificate for both the filter and filter holder.

- Before the start of sterilization process the autoclave chamber leak test should be automatically launched; after test completion the printer should print out the report on its positive result.

Desirable siting: